## Evaluation Tool
Cat or SubCat Code: | For example: 18-823 | BMET completes Green Sections | BME Asset Tag Req'd: | YES
UMDNS Description: | For Example: Ultrasonic Therapy Systems
Evaluation Criteria: Risk evaluation method based on World Health Organization guidelines. |  |  | Scheduled Maintenance Req'd: ↓
Equipment Function Therapeutic, Diagnostic, Analytical and Miscellaneous equipment categories are considered. Note: This score in NOT based on equipment location (e.g. ICU) rather you are being asked whether the equipment's FUNCTION is Therapeutic, Diagnostic or Analytical regardless of it's physical location. For example a thermometer has a diagnostic score of 6, even though it may be located in a critical care area, while a defibrillator has a score of 10 even if located in a general ward. |  | 9 | Scheduled Maintenance Required: | YES
Physical Risk Associated with Clinical Application Lists the potential patient or equipment risk during use. |  | 3 | Priority: | NORMAL
Scheduled Maintenance Requirement Describes the level and frequency of "scheduled maintenance" required as noted by the Manufacturer or through experience. |  | 3 | Frequency: | 1/YEAR
Equipment History Any information available regarding service history that should be considered (repair frequency, patient incidents, device alerts) |  | 0 | Note: The above calculated result determines whether an equipment category/sub-category is to be included as Scheduled Maintenance in the BC BME CMMS. Unscheduled Maintenance occurs after equipment repairs &/or as determined by local BME supervisors.
Request a review: If BMET disagrees with the calculated Scheduled Maintenance Result; request a review by completing the Review Tab (below) and put an R in this box-------------->
Brief Comment: (optional): |  |  | After completing this form
Health Authority Mandated (This field is set by the reviewing BME Risk Mgr/Engineer) Purpose: To overide the scored Priority & Frequency. (typically used to increase the Priority &/or Frequency due to a Corporate Directive or unique local need where a BMET Requested a Review. |  | 0 | · Save as: <UMDNS Code>.xls on your computer. · Email completed forms to: DBMTeam@bcbiomed.org
Input: Equip. Function SCORE |  | 9
Input: Physical Risk SCORE |  | 3
Input: Sched Maint. Required SCORE |  | 3
Input: Equip. History SCORE |  | 0
Input: Review Requested? (Added feature) * |  | 0
Input: Mandated ? (Added feature) * |  | 0
Input: Override by Review TAB - Priority |  | Review
Input: Override by Review TAB - Frequency |  | Review
Calc: EM # Asset Tag (SUM of Funct + Risk + Maint) |  | 15
Calc: EM# Sched. Maint. (SUM of Funct + Risk + Maint + History) |  | 15
If: EM # Asset Tag > 10 (WHO is 12 but then only PM'd are Inventoried) * |  | YES
If: EM # Sched. Maint > 12 |  | YES
If: EM # Sched. Maint ≥ 16 (WHO is 17) * |  | NO
If: Maint Required > 3 |  | NO
If: Maint Required ≤ 3 |  | YES
If: Equip History < 0 |  | NO
If: Override Requested |  | NO
If: Mandated > 0 |  | NO
Logic: OR #1 (PM Required Indicated) |  | TRUE
Logic: AND # 1 (Determine Criticality) |  | TRUE
Logic: AND # 2 (2/Year Indicated) (4 month negated) * |  | FALSE
Logic: AND # 3 (1/Year Indicated) |  | TRUE
Logic: AND # 4 (24 Month Indicated) (24 Month added) * |  | FALSE
Logic Override:
10
9
8
7
6
5
4
3
2
5
4
3
2
1
2
1
0
-1
-2
1 | Mandatory or Unique Frequency & Priority as Established on Review Tab
0
Yes
No
R
Note on C4: Biomedical Engineering: Enter the model number and a common name for the device being evaluated. For example: 2M9163 Inf Pump
Note on D6: 10 = Therapeutic Life Support 9 = Therapeutic Surgical & Intensive Care 8 = Therapeutic Physical Therapy & Treatment 7 = Diagnostic Surgical & Intensive Care Physiological Monitoring & Ionizing radiation emitting devices (e.g. x-ray) 6 = Diagnostic Additional Physiological Monitoring, Diagnostics & non-ionizing types of medical imaging 5 = Analytical Analytical Laboratory equipment 4 = Analytical Laboratory accessories 3 = Analytical Computers & related 2 = Misc. Patient related & other (e.g. shavers, cast cutters)
Note on D7: 5 = Death 4 = Patient or Operator Injury 3 = Misdiagnosis or inappropriate therapy 2 = Equipment damage 1 = No significant or identifiable risks
Note on D8: 5 = Extensive 4 = Above Average 3 = Average (approximately 1/year) 2 = Below Average 1 = Minimal (would typically replace rather than servce. e.g.) inexpensive
Note on D9: +2 = Significant (more than one incident every 6 months) +1 = Moderate ( one incident every 6 -9 months) 0 = Average (one incident every 9 - 18 months) -1 = Low (one incident every 18 - 30 months) -2 = Insignificant (Less than 1 incident per 30 months)
Note on B11: Biomedical Engineering: e.g.1) Devices processes Blood & Blood Products e.g.2) Device on Service Contract
Note on E11: Biomedical Engineering: If the Model number is too long or contains special characters just mention the model name in your email.
Note on D12: 1 = Mandatory or Unique Frequency as established on Review tab (e.g. Service contracts) 0 = For All Other

## Review
Scheduled Maintenance Review Form | BMET completes Green sections (Note most fields have drop down boxes)
Is this a request to over-ride a BC CEC 2006 Standard? If YES, answer all the following:
1) For which Health Authority (choose one):
2) For Site - Type Full Site Name e.g. Royal Inland (not RIH):
3) Your suggested Scheduled Maintenance Priority:
4) Your suggested frequency for Scheduled Maintenance:
If Other Frequency describe here:
5) Select the most appropriate reason for over- riding BC CEC 2006 standard:
If Other reason describe here:
6) Is Scheduled Maintenance Mandated by an External Accrediting Agency?
If yes, please indicate the agency:
7) Likelihood Scheduled Maintenance will improve safety of device relative to applicable standards. (e.g. Health Canada, CSA, ANSI, ISO etc.)
8) What is the probability or potential the device will function improperly if not provided with suggested level of Scheduled Maintenance?
9) What is the risk to the Patient, Operator, or Regional service if the device does not function properly?
This section for For Manager/Engineer (if approving an over-ride) | Note: Values entered in fields C16 & C17 (below) will overwrite the Evaluation Tool's Calculations - once the Approver sets (D12) on Evaluation Tab > 0
Scheduled Maintenance Priority: | Review
Scheduled Maintenance Frequency: | Review
Approved by: (Name, Title, Date)
Approver's Additional Notes:
Yes
No
All Health Authorities
VIHA
VCH
PHC
PHSA
IH
NH
Critical
Normal
No Scheduled PM Req'd
2/Year
1/Year
24 Month
Other
None
1. CMMS evidence is supporting change in frequency of current Scheduled Maintenance
2. Device design is driving change in frequency of Scheduled Maintenance
3. Specific patient incident is driving change in frequency of Scheduled Maintenance
4. How clinicians use the device is driving change in frequency of Scheduled Maintenance
5. Unique environment driving change in frequency (e.g. very low utilization or device used in patient's home)
6. New evidence is supporting change in frequency of Scheduled Maintenance (please note source below in "other")
7. Patient population driving Priority and/or Frequency change.
8. Other
Extreme
Very High
High
Moderate
Low
Very Low
Not Applicable
Note on C16: Biomedical Engineering: Upon completion of review enter: Critical or Normal or None
Note on C17: Biomedical Engineering: Upon completion of review enter frequency using format: X/Year X Months X Weeks X Days

## Tool Logic
Each green hexagon is scored during incoming inspection by the BMET. Scores are summed to generate a two Equipment Maintenance Numbers (EM#). If the Asset Tag EM# is greater than 10 the device qualifies for inclusion in the Computerized Maintenance Management System (CMMS). If the Schedule Maintenance EM# is greater than 12 the device qualifies for Schedule Maintenance (SM). If SM EM# is greater than 16, Maintenance is Critical, scores of 13-15 are considerde NORMAL, scores of 12 or less indicate the device does not require SM. Critical devices are scheduled for maintenance 2/year if the Required Maintenance score is above average. NORMAL devices are scheduled for either annual or semi-annual maintenance as determined by their Required Maintenance & History scores. Over-rides can be generated to account for Mandated or Unique requirements which do not fit the normal model.

## For Export to BC PM Sched. (hidden)
Fct'l Group | Equip. Model, Common Name | EM# | PM Req'd | PM Ranking | PM Freq. | Comments
| For Example: Ultrasonic Therapy Systems | 3 | YES | NORMAL | 1/YEAR
Copy Cells C3 through H3; then on the BC PM Schedule Use: Paste Special <Values>
Note on A2: Biomedical Engineering: Clinical Imaging Lab
Note on C2: Examples: 2M9163 Inf. Pump
Note on H2: Biomedical Engineering e.g.) Mandated - CSA Std.

## Executive Summary
Executive Summary
Title |  | Scheduled Maintenance Evaluation Tool
Purpose |  | To provide a method for determining whether a medical device should be included in the Scheduled Preventive Maintenance program.
Top Risks | 1) 2) 3) | Decisions on Scheduled Maintenance not linked to a recognized standard. (Accreditation) Lack of Standardization between Health Authorities. Computerized Maintenance Management System (CMMS) not adequately maintained.
Lead |  | Greg Brett, Manager Biomedical Engineering Risk & Quality, IHA
Sponsor |  | Tim Rode, Director Biomedical Engineering, IHA
Recommendation
| 1) 2) 3) | That IH Biomedical Engineering adopt a Preventive Maintenance Scheduling method based upon the recommendations of the World Health Organization. That IH Biomedical Engineering present the scheduling method to the BCCEC, seek constructive feedback/collaboration and discuss possible implementation across all B.C. Health Authorities. That a tool be developed to aid BMETs with evaluating new medical device's for inclusion in scheduled preventative maintenance.
Background
Presently, a medical device's inclusion in the scheduled preventive maintenance program is determined by referencing a document agreed to by all BC CEC members in 2006. New types of medical devices are constantly being introduced and opportunities to update the 2006 document have proven to be infrequent. As a result, BMETs in the field may not have sufficient guidance as to PM requirements. Medical devices not listed in the original 2006 BC CEC document, continue to be evaluated for Scheduled PM based on device Risk & PM Utility at a local level. There is some subjectivity involved in this scoring process which can lead to inconsistent scoring between Health Authorities. Increasingly, Biomedical Engineering Departments are required to report their Schedule Preventive Maintenance compliance and meet specific targets for Hospital Accreditation. Without standardization in Scheduled PM Frequency and Criticality accurate reports and the ability to compare Provincial norms is greatly compromised.
Discussion
· The Evaluation Tool attached to this document is a simple spreadsheet, which builds upon the foundational concepts found in the World Health Organizations (WHO) Medical Equipment Maintenance Overview. The intent of the tool is to provide an efficient evaluation method which could be made available to all BMETS, while still yielding an outcomes similar to those in the original CEC2006 document. · Separate schedules for Safety Inspections & Preventative Maintenance are not generated as both of these factors are considered during every Scheduled PM event. · The attached Tool sets specific parameters for inclusion in the Scheduled Preventive Maintenance program. (Inclusion: Yes/No; Criticality Critical/Normal; and Frequency 2/Year, 1/Year or 24 months) · The Tool is designed to be used to evaluate specific models of medical devices regardless of their physical location in the hospital. (i.e. Utilization is not currently factored in) This decision was made because utilization rates change with time and tracking this requires considerable additional workload, with questionable productivity gains. · The Tool adds one additional frequency to the WHO plan. This 24 month frequency is implemented where the known equipment history has demonstrated that an extended frequency is a practical & safe decision. ·
References
World Health Organization, Medical Equipment Maintenance Program Overview
The WHO document references the following authors:

## BC CEC 2006
Mandatory Inspections | Recommended Inspections